## Slide 1
Low energy section design of a linear accelerator for carbon ions therapy
V. Bencini

## Slide 2
Outline
Introduction
  Hadrontherapy
  Accelerators for hadrontherapy
  The linac solution
TwinEBIS source
  Working principles
  Output parameters
LEBT
  General layout
  Matching
  Error study
The 750 MHz RFQ
  Input parameters
  Final design
  End to end tracking
Conclusions and next steps

## Slide 3
Introduction

## Slide 4
Hadron therapy

## Slide 5
Hadron therapy
Tumore

## Slide 6
Accelerators for hadron therapy
Features
Small footprint
Drawbacks
Passive energy modulation
Beam emittance
Slow energy modulation
Only for protons

## Slide 7
Accelerators for hadron therapy
Cyclotrons
Features
Small footprint
Drawbacks
Passive energy modulation
Beam emittance
Slow energy modulation
Only for protons

## Slide 8
Accelerators for hadron therapy
Cyclotrons
Features
Small footprint
Drawbacks
Passive energy modulation
Beam emittance
Slow energy modulation
Only for protons
Synchrotrons
Features
Active energy modulation
Both protons and carbon ions
Drawbacks
Slow energy change
Big footprint
Big emittance

## Slide 9
Accelerators for hadron therapy
Cyclotrons
Features
Small footprint
Drawbacks
Passive energy modulation
Beam emittance
Slow energy modulation
Only for protons
Synchrotrons
Features
Active energy modulation
Both protons and carbon ions
Drawbacks
Slow energy change
Big footprint
Big emittance
Features
Active energy modulation
Fast energy modulation
Both protons and carbon ions
Drawbacks
Bigger than cyclotron
‘linear’ footprint
Linacs

## Slide 10
The linac solution
100 MeV/u
0.015 MeV/u
2.5/5 MeV/u
100 MeV/u
Energy modulation
Fixed energy
RFQ
TwinEBIS
Patient
All klystrons off
100 MeV/u
Parameter | Value
Frequency | 3 GHz
Species | 12C6+
Current required (average) | 0.8 nA
Final energy | 100-430 MeV/u
Repetition rate | 200 Hz

## Slide 11
The linac solution
100 MeV/u
0.015 MeV/u
2.5/5 MeV/u
100 MeV/u
Fixed energy
RFQ
TwinEBIS
Patient
300 MeV/u
Some klystrons on
300 MeV/u
Parameter | Value
Frequency | 3 GHz
Species | 12C6+
Current required (average) | 0.8 nA
Final energy | 100-430 MeV/u
Repetition rate | 200 Hz

## Slide 12
The linac solution
150 MeV/u
0.015 MeV/u
2.5/5 MeV/u
100 MeV/u
Energy modulation
Fixed energy
RFQ
TwinEBIS
Patient
430 MeV/u
All klystrons on
Parameter | Value
Frequency | 3 GHz
Species | 12C6+
Current required (average) | 0.8 nA
Final energy | 100-430 MeV/u
Repetition rate | 200 Hz

## Slide 13
The linac solution
150 MeV/u
0.015 MeV/u
2.5/5 MeV/u
100 MeV/u
Energy modulation
Fixed energy
RFQ
TwinEBIS
Patient
All klystron on
430 MeV/u
One design already proposed (CABOTO - Fondazione TERA)
Next step is to propose an alternative design
Parameter | Value
Frequency | 3 GHz
Species | 12C6+
Current required (average) | 0.8 nA
Final energy | 100-430 MeV/u
Repetition rate | 200 Hz

## Slide 14
The linac solution
150 MeV/u
0.015 MeV/u
2.5/5 MeV/u
100 MeV/u
Energy modulation
Fixed energy
RFQ
TwinEBIS
Patient
All klystron on
430 MeV/u
Parameter | Value
Frequency | 3 GHz
Species | 12C6+
Current required (average) | 0.8 nA
Final energy | 100-430 MeV/u
Repetition rate | 200 Hz
The work here presented will be focused on the low energy section

## Slide 15
Low Energy Beam Transport (LEBT)
Electrostatic LEBT
Transports and matches the beam into the RFQ acceptance
TwinEBIS
Electron Beam Ion Source
Pulsed source
Produces 1·109 12C6+ ions in a 5 us pulse (~200 times higher than requirements)
RFQ
First RF accelerating structure
Acceleration, bunching, focusing
Boost particles to 2.5/5 MeV/u (two designs)
The linac solution – low energy section

## Slide 16
The TwinEBIS source

## Slide 17
Working principle
How it works?
Courtesy of M. Breitenfeldt

## Slide 18
Working principle
MEDeGUN
How it works?
Charge breeding by ionization with electron beam
Ions kept in a potential trap
Open the trap and extraction of the ions
Courtesy of M. Breitenfeldt

## Slide 19
Output parameters
MEDeGUN
Courtesy of M. Breitenfeldt
How it works?
Charge breeding by ionization with electron beam
Ions kept in a potential trap
Open the trap and extraction of the ions
Design parameter | MEDeGUN
Test site | TwinEBIS, CERN
Main magnet | 2 T
Trap length | 0.25 m
Electron current | 1 A
Current density | 1.5 kA/cm2
Electron energy | 7.5-10 keV
Capacity C6+ | 1·109 Ions per pulse
Repetition rate | 200 Hz
Current | 160 nA

## Slide 20
The LEBT

## Slide 21
General Layout
Electron collector
Extraction electrode
Adaptor electrode
Acceleration gap
Gridded lenses
Einzel lens
Angle deflectors
Switchyard deflector
RFQ port
Additional ports
Faraday cups
Pepperpot
Extract ions at the energy of 6.5 keV/u and accelerate them up to 15 keV/u

## Slide 22
General Layout
Electron collector
Extraction electrode
Adaptor electrode
Acceleration gap
Gridded lenses
Einzel lens
Angle deflectors
Switchyard deflector
RFQ port
Additional ports
Faraday cups
Pepperpot
Extract ions at the energy of 6.5 keV/u and accelerate them up to 15 keV/u
Focusing

## Slide 23
Matching methodology
Step 1 – Telescopic matching
Select voltage on adaptor that gives alpha=0
Select voltage on gridded lens 1 that gives alpha=0
Select voltage on Einzel lens that gives alpha=0
Scan gridded lens to to find best match to accptance
Step 2 – Optimised matching
Use Telescopic matching set-up as starting point
Scan many possible combination of voltages
Obtain transmission maps
Fine scan on gridded lens 2

## Slide 24
The LEBT – results
k
l
m
n
o
n
Beam current | 0 mA | 0.38 | 3 mA
Particles into acceptance | 99% | 96% | 88%
q
Three beam currents were considered: 0 mA (no space charge case), 0.38 mA (current expected from present set up) and 3 mA (max theoretical current form TwinEBIS).

## Slide 25
The LEBT – error study
For each emittance value scan Twiss alpha and beta  generate a space of input distributions
Each beam matched to the RFQ acceptance (considering 0.34 mA current)
Select the set of input beam that could be matched into RFQ acceptance with efficiency higher than 90%
Still work in progress

## Slide 26
Radio Frequency Quadrupole

## Slide 27
Radio Frequency Quadrupole
RF resonator in TE210 mode (quadrupole)
Modulation of the vane profile induce longitudinal field  ACCELERATION
RF provides temporal structure to the beam

## Slide 28
Input parameters
Input energy defined by the LEBT.
The choice of design aperture and vane voltage is a compromise between power consumption () and RFQ acceptance (increasing with aperture).
The initial phase was set to -50: only the particles that fit in the longitudinal bucket of the structure following the RFQ are captured.
All the losses occur at low energy, avoiding activation of the cavities.

## Slide 29
Final design
2.5 MeV/u
5 MeV/u

## Slide 30
Final design
2.5 MeV/u
5 MeV/u
Transmission low on purpose

## Slide 31
End to end tracking
Parameter | IH-structure | DTL-structure
Input energy | 2.5 MeV/u | 5 MeV/u
Accelerating gradient | 5.7 MV/m | 10 MV/m
Sync. Phase | -12 | -30
RF frequency | 750 MHz | 3 GHz
Parameter | IH-structure | DTL-structure
Input energy | 2.5 MeV/u | 5 MeV/u
Accelerating gradient | 5.7 MV/m | 10 MV/m
Sync. Phase |  | 
RF frequency | 750 MHz | 3 GHz
Two structure could follow the RFQ (two different longitudinal acceptances)

## Slide 32
Conclusions and next steps

## Slide 33
Conclusions
LEBT matching was performed at different current and the maximum error on th einput distribution was defined
Tracking through the RFQ shows the capability of the system to provide an ion current that exceeds significantly the clinical requirements.
The EBIS+LEBT+RFQ system would be the perfect injector for a linac for carbon ion therapy.
Next steps
The LEBT will be installed and tested during the next year. Installation, commissioning and measurements will be part of the thesis work.
Collaboration with RF group to select the accelerating structure of the all linac
Beam dynamics design of the linac
Conclusions and next steps

## Slide 34
Thank you for your attention!